HTIC FETAL ULTRASOUND
ALL BENCHMARKS & EXPERIMENTS

Consolidated technical record of dataset checks, model quality, optimization, ONNX conversion, CPU performance and pipeline profiling

Project root (Windows): E:\HTIC_FETAL_ULTRASOUND

Purpose: bring the measured results and experiment/report inventory into one structured reference. This document records what was measured, what was not measured, and which saved files contain supporting evidence.

# 1. Executive Summary

Dataset: FETAL_PLANES_DB, 12,400 CSV-referenced images across 1,792 patients and six classes.

Patient-level split: Train=1 development pool (7,129 images / 896 patients) and Train=0 held-out test (5,271 images / 896 patients); zero patient overlap. Development pool split into 5,673 training and 1,456 validation images using StratifiedGroupKFold (5 folds, shuffle=True, random_state=42).

Baseline: pretrained MobileNetV3-Large fine-tuned for six classes. Held-out test accuracy 92.1647%, macro-F1 91.1170%.

Training-time optimization: AMP / FP16 mixed precision training was completed. Final validation accuracy 93.96%, macro-F1 93.17%; training duration 430.99 seconds and peak GPU memory 786.2 MB. A comparable baseline training-time measurement was not captured, so no measured training speedup is claimed.

Deployment candidate: FP16 ONNX. File size 8.0548 MiB vs FP32 16.0493 MiB (49.81% smaller). Validation accuracy unchanged at 94.0934%; macro-F1 decreased from 93.23095% to 93.16966%. Predictions differed on 4/1,456 validation images (99.7253% agreement).

Laptop CPU model-only benchmark: FP16 median latency 4.4083 ms vs FP32 4.6826 ms, a 5.86% reduction; throughput 225.18 vs 197.28 inferences/s (+14.14%). This was one benchmark run using synthetic preprocessed input, not full image-to-prediction.

CPU thread tuning: controlled 8-thread vs 1-thread pipeline profiling reduced median end-to-end latency from 25.840 ms to 17.423 ms (32.6%).

Batching: batch 4 reached 277.72 images/s vs 213.76 images/s at batch 1 (+29.9%) on preprocessed inputs. This is a throughput result and does not include queueing delay.

INT8 dynamic and static quantization substantially degraded validation classification quality; they were retained as experiments but rejected for deployment.

Laptop GPU inference was not measured: the installed ONNX Runtime exposes CPU only; GPU package installation was cancelled. Jetson and tablet performance were not measured; a separate proposal describes how to test them.

# 2. Dataset and Evaluation Integrity

## 2.1 Dataset audit

| Item | Recorded result |
| --- | --- |
| Dataset | FETAL_PLANES_DB |
| CSV rows / referenced images | 12,400 |
| Image files in data/Images | 12,401 PNG files; one unreferenced duplicate-looking “(1)” image was noted |
| Patients | 1,792 |
| Classes | 6 |
| Missing referenced images | None reported |
| Image formats | Grayscale and RGBA images encountered; preprocessing converts through the official pretrained transform pipeline |

| Class | Image count |
| --- | --- |
| Other | 4,213 |
| Fetal brain | 3,092 |
| Fetal thorax | 1,718 |
| Maternal cervix | 1,626 |
| Fetal femur | 1,040 |
| Fetal abdomen | 711 |

## 2.2 Patient-isolated data partitions

| Partition | Images | Patients | Purpose |
| --- | --- | --- | --- |
| Development pool (CSV Train=1) | 7,129 | 896 | Training and validation only |
| Training subset | 5,673 | Not separately recorded here | Model fitting |
| Validation subset | 1,456 | Not separately recorded here | Model selection and conversion-quality checks |
| Held-out test (CSV Train=0) | 5,271 | 896 | Final baseline generalization evaluation |

Training/validation split method: StratifiedGroupKFold(n_splits=5, shuffle=True, random_state=42).

The held-out test patient IDs do not overlap the development-pool patient IDs.

Important: FP32 baseline held-out test metrics are available. FP16 held-out test metrics had not yet been run at the time this consolidated document was prepared. Validation metrics must not be presented as test metrics.

# 3. Baseline Model and Training

## 3.1 Model and setup

| Setting | Value |
| --- | --- |
| Architecture | MobileNetV3-Large |
| Initialization | MobileNet_V3_Large_Weights.DEFAULT pretrained weights |
| Classifier output | 6 classes |
| Training framework | PyTorch 2.11.0+cu130; torchvision 0.26.0+cu130 |
| Training hardware | Google Colab Tesla T4 GPU |
| Optimizer | AdamW |
| Learning rate | 1e-4 |
| Loss | Cross-entropy |
| Batch size / epochs | 32 / 5 |
| Baseline checkpoint | artifacts/baseline/best_model.pth |
| Preprocessing | MobileNet_V3_Large_Weights.DEFAULT.transforms() |

## 3.2 Baseline validation history

| Epoch | Validation accuracy | Validation macro-F1 |
| --- | --- | --- |
| 1 | 88.19% | 85.54% |
| 2 | 86.06% | 80.97% |
| 3 | 89.35% | 85.24% |
| 4 | 93.48% | 92.40% |
| 5 | 94.09% | 93.23% |

## 3.3 Held-out test baseline

| Metric | Result |
| --- | --- |
| Accuracy | 92.1647% |
| Macro-F1 | 91.1170% |
| Test images | 5,271 |

The held-out test set was patient-isolated from the development pool. These values are the reference test results for comparison. Do not compare them directly with validation results as if they were the same split.

# 4. Training-Time Optimization: AMP / Mixed Precision

AMP training used FP16 autocast with GradScaler for training; validation was evaluated in FP32. This is distinct from converting a trained FP32 ONNX model to FP16.

| Epoch | AMP validation accuracy | AMP validation macro-F1 |
| --- | --- | --- |
| 1 | 84.75% | 80.42% |
| 2 | 88.87% | 85.16% |
| 3 | 91.69% | 88.94% |
| 4 | 93.34% | 91.89% |
| 5 | 93.96% | 93.17% |

| Measurement | Recorded value |
| --- | --- |
| Total AMP training duration | 430.99 seconds |
| Peak GPU memory | 786.2 MB |
| Final AMP validation accuracy | 93.96% |
| Final AMP validation macro-F1 | 93.17% |

Interpretation: final AMP validation quality was close to the baseline validation result.

Limitation: a comparable baseline training duration and peak-memory record was not captured. Therefore the record does not support a numerical AMP training-speedup or memory-reduction claim.

Supporting files: src/train_mixed_precision.py; artifacts/experiments/amp_fp16/; reports/amp_fp32_comparison.json; reports/amp_fp32_comparison.md.

# 5. ONNX Export and Numerical Verification

| Artifact / check | Recorded result |
| --- | --- |
| FP32 ONNX | artifacts/onnx/mobilenetv3_fetal_fp32.onnx |
| FP32 file size | 16.05 MiB (benchmark JSON reports 16.0493 MiB) |
| FP32 input/output | Input [1,3,224,224], output [1,6] |
| Graph | Opset 17; 140 nodes; fixed batch |
| Dynamic-batch FP32 | artifacts/onnx/mobilenetv3_fetal_fp32_dynamic_batch.onnx |
| Dynamic input | ['batch_size',3,224,224] |
| PyTorch-vs-ONNX smoke test | Maximum absolute logit difference 0.0000124; same argmax on one random input |

The PyTorch-vs-ONNX check is a one-input smoke test, not a dataset-wide accuracy evaluation. Verification artifacts are in reports/onnx_verification/verification_report.txt and verification_results.json.

# 6. Post-Training Optimization: FP16 ONNX

## 6.1 Conversion and artifact

Converted the FP32 ONNX graph to FP16 using onnxconverter_common.float16.convert_float_to_float16(..., keep_io_types=True).

Primary deployment candidate: artifacts/onnx/mobilenetv3_fetal_fp16.onnx.

Experiment copy retained at artifacts/experiments/onnx_fp16/mobilenetv3_fetal_fp16.onnx; the duplicate is intentionally retained for provenance.

FP16 graph passed ONNX graph checking. Conversion emitted warnings that very small values were clipped to ±1e-7; the graph remained loadable and was evaluated.

The external model input/output interface remains float32; internal graph operations use FP16 where converted.

## 6.2 FP32 versus FP16 validation quality

| Metric | FP32 ONNX | FP16 ONNX | Difference / note |
| --- | --- | --- | --- |
| Accuracy | 94.0934% | 94.0934% | No accuracy change |
| Macro-F1 | 93.23095% | 93.16966% | −0.0613 percentage points |
| Correct predictions | 1,370 / 1,456 | 1,370 / 1,456 | Same correct count |
| Prediction agreement | — | 99.7253% | 4 of 1,456 predictions differed |

Validation result file: reports/fp32_vs_fp16_validation.json. This is validation-set evidence, not held-out test evidence.

## FP32 versus FP16 validation quality

| Metric | FP32 Baseline | FP16 Optimized |
| --- | --- | --- |
| Held-out test accuracy | 92.1647 % | 92.1836 % |
| Held-out test macro-F1 | 91.1170 % | 91.1595 % |
| ONNX model size | 16.05 MiB | 8.05 MiB |
| Model size reduction | — | 49.81 % |

## 6.3 Laptop CPU FP32 versus FP16 benchmark

Environment: Windows laptop; ONNX Runtime 1.31.0, CPUExecutionProvider, 8 intra-op threads; 30 warm-up calls and 200 measured calls; same deterministic synthetic input shape [1,3,224,224]. Model-only timing excludes image decoding and preprocessing. Values below are from a single recorded run.

| Metric | FP32 ONNX | FP16 ONNX | FP16 vs FP32 |
| --- | --- | --- | --- |
| Model size | 16.0493 MiB | 8.0548 MiB | −49.81% |
| Median latency | 4.6826 ms | 4.4083 ms | −5.86% |
| P95 latency | 7.3332 ms | 4.9821 ms | Lower in this run |
| Mean latency | 5.0691 ms | 4.4409 ms | −12.39% |
| Throughput | 197.28 inferences/s | 225.18 inferences/s | +14.14% |

Interpretation: FP16 reduced artifact size by about half and was modestly faster in this model-only CPU run.

Caveat: one run is not a robust statistical estimate; synthetic input does not represent the full image-to-prediction pipeline. Keep this result separate from full pipeline timing.

Supporting files: src/benchmark_fp32_fp16.py; benchmarks/fp32_vs_fp16_laptop.json.

# 7. INT8 Quantization Experiments (Rejected for Deployment)

## 7.1 Dynamic INT8

| Metric | Recorded result |
| --- | --- |
| Artifact | artifacts/experiments/onnx_dynamic_int8/mobilenetv3_fetal_dynamic_int8.onnx |
| Model size | Approximately 4.22 MiB |
| Size reduction vs FP32 | Approximately 73.7% |
| Validation accuracy | 72.12% |
| Validation macro-F1 | 65.17% |

The validation-quality reduction was too large for the intended classification task. Dynamic INT8 was not selected for deployment. Detailed latency values are preserved in benchmarks/fp32_vs_dynamic_int8_cpu.json;

| Metric | FP32 ONNX | Dynamic INT8 ONNX |
| --- | --- | --- |
| Mean latency (ms/image) | 32.764 ms | 160.806 ms |
| Median latency (ms/image) | 34.353 ms | 131.420 ms |
| Throughput (images/second) | 30.52 | 6.22 |
| Relative speed | 1× (baseline) | 4.91× slower |

Measured CPU performance (ONNX Runtime CPUExecutionProvider): dynamic INT8 mean latency was 160.806 ms versus 32.764 ms for FP32 (4.91× slower); median latency was 131.420 ms versus 34.353 ms; throughput was 6.22 images/s versus 30.52 images/s. These measurements, alongside the validation-quality loss, support rejecting dynamic INT8 for this CPU deployment. The observed slowdown is specific to this model, runtime, CPU and benchmark setup; do not generalize it to all CPUs.

## 7.2 Static INT8

| Variant | Model size | Validation accuracy | Validation macro-F1 |
| --- | --- | --- | --- |
| Full-range static INT8 | Approximately 4.50 MiB | 51.58% | 27.75% |
| Reduced-range static INT8 | Approximately 4.50 MiB | 56.52% | 34.24% |

Calibration set: 204 development-pool images, 34 per class, seed 42.

Static INT8 artifacts include artifacts/experiments/onnx_static_int8/mobilenetv3_fetal_preprocessed.onnx and mobilenetv3_fetal_static_int8.onnx; calibration manifest and metadata are stored in the same folder.

Even reduced-range static INT8 had unacceptable validation quality. It was rejected for deployment.

Supporting files: benchmarks/fp32_vs_static_int8_cpu.json; reports/quantization_validation_results.json; reports/quantization_decision.md; reports/dynamic_int8_findings.md.

No quantization-aware training (QAT) was performed; the user explicitly chose not to pursue QAT or further quantization experiments.

# 8. CPU Thread-Count Benchmark

Script: src/benchmark_onnx.py. Model-only inference with synthetic input; 30 warm-up calls and 200 measured calls. ONNX Runtime CPUExecutionProvider.

| CPU threads | Median latency | P95 latency | Mean latency | Throughput |
| --- | --- | --- | --- | --- |
| 1 | 12.242 ms | 15.454 ms | 12.712 ms | 78.66 inferences/s |
| 2 | 6.965 ms | 11.005 ms | 7.627 ms | 131.12 inferences/s |
| 4 | 5.736 ms | 6.744 ms | 5.878 ms | 170.12 inferences/s |
| 8 | 4.178 ms | 4.874 ms | 4.271 ms | 234.14 inferences/s |

Eight threads were fastest among the tested settings for this model-only workload.

Supporting files: benchmarks/laptop_thread_benchmark.json; benchmarks/onnx_cpu_comparison.json. These JSONs may preserve additional fields and should remain the source of record.

# 9. Batch-Size Benchmark

Script: src/benchmark_batch_size.py. Dynamic-batch FP32 ONNX model; 128 preprocessed images; 8 CPU threads; 10 warm-up calls and 30 repetitions.

| Batch size | Median ms/image | Throughput | Batch-call latency |
| --- | --- | --- | --- |
| 1 | 4.678 ms | 213.76 images/s | 4.557 ms |
| 4 | 3.601 ms | 277.72 images/s | 14.284 ms |
| 8 | 3.622 ms | 276.11 images/s | 28.773 ms |
| 16 | 3.770 ms | 265.28 images/s | 60.030 ms |

Batch size 4 achieved the highest measured throughput: approximately +29.9% vs batch size 1.

Batching improves throughput for queued/multiple images, but a batch of 4 takes longer to return than a single-image call. It may add queueing delay and is not automatically the best setting for interactive single-image inference.

Supporting file: benchmarks/batch_benchmark.json.

# 10. End-to-End Pipeline Profiling

These measurements are CPU laptop profiling results using ONNX Runtime and the official image transform. Stage timings are kept distinct from model-only synthetic-input benchmarks. The profiles use different scripts, image sets, warm-up counts and repetitions, so they should not be combined into one table as if they were a single controlled experiment.

## 10.1 Earlier one-image stage profile

| Stage | Mean | Median | P95 |
| --- | --- | --- | --- |
| Image loading | 8.180 ms | 8.044 ms | 8.872 ms |
| Preprocessing | 6.633 ms | 6.217 ms | 9.298 ms |
| Inference | 4.958 ms | 4.838 ms | 5.541 ms |
| End-to-end | 19.780 ms | 19.270 ms | 23.569 ms |

Script: src/profile_inference_stages.py. Earlier profile; use as exploratory evidence rather than the final controlled comparison.

## 10.2 Controlled pipeline profile: 1 versus 8 threads

Script: src/profile_pipeline.py. Same controlled methodology: 50 sorted unique images, 2 warm-up rounds × 5 images, 5 measurement rounds; CPUExecutionProvider; official transforms.

| Stage / metric | 1 thread | 8 threads | Change at 8 threads |
| --- | --- | --- | --- |
| Loading — mean | 6.213 ms | 6.749 ms | Slightly slower mean |
| Loading — median | 5.906 ms | 6.518 ms | Slightly slower median |
| Loading — P95 | 10.169 ms | 10.279 ms | Similar |
| Preprocessing — mean | 5.808 ms | 6.125 ms | Similar/slightly slower |
| Preprocessing — median | 5.497 ms | 5.724 ms | Similar/slightly slower |
| Preprocessing — P95 | 6.319 ms | 7.488 ms | Higher in this run |
| Inference — mean | 14.487 ms | 5.140 ms | −64.5% |
| Inference — median | 14.822 ms | 4.915 ms | −66.8% |
| Inference — P95 | 16.061 ms | 5.923 ms | −63.1% |
| End-to-end — mean | 26.515 ms | 18.022 ms | −32.0% |
| End-to-end — median | 25.840 ms | 17.423 ms | −32.6% |
| End-to-end — P95 | 32.044 ms | 22.694 ms | −29.2% |

Interpretation: most of the gain came from model inference; loading and preprocessing did not improve. This is a meaningful code/runtime configuration optimization because thread count was tuned against measured latency.

Supporting files: benchmarks/pipeline_profile_controlled_1threads.json, pipeline_profile_controlled_4threads.json, pipeline_profile_controlled_8threads.json, and pipeline_profile_controlled.json (older controlled-profile artifact).

## 10.3 Extended 8-thread pipeline profile

| Stage / metric | Mean | Median | P95 |
| --- | --- | --- | --- |
| Image loading | 6.502 ms | 6.248 ms | 10.211 ms |
| Preprocessing | 5.991 ms | 5.481 ms | 7.099 ms |
| Inference | 4.845 ms | 4.612 ms | 5.408 ms |
| Postprocessing | 0.114 ms | 0.111 ms | 0.135 ms |
| End-to-end | 17.460 ms | 16.586 ms | 21.944 ms |

Script: src/profile_pipeline_extended.py. 50 fixed images, 10 warm-ups, 5 rounds, 8 threads. Process RSS was sampled: before 321.27 MB, observed sampled peak 340.65 MB. This is a sampled process-RSS observation, not a guaranteed absolute peak.

One example prediction was class index 0 (Other), confidence 94.0156%. This is an example output, not an accuracy measure.

## 10.4 Repeated single-image pipeline profile

| Stage / metric | Mean | Median | P95 |
| --- | --- | --- | --- |
| Image loading | 11.416 ms | 11.482 ms | 12.368 ms |
| Preprocessing | 6.249 ms | 6.190 ms | 7.082 ms |
| Inference | 4.722 ms | 4.669 ms | 5.140 ms |
| Postprocessing | 0.114 ms | 0.111 ms | 0.141 ms |
| End-to-end | 22.509 ms | 22.623 ms | 24.233 ms |

Script: src/benchmark_repeated_prediction.py. One fixed image, 20 warm-up calls and 100 measured calls, 8 threads. Disk/image loading was measured repeatedly; this result is not directly comparable to the 50-image profile because workload and cache conditions differ.

## 10.5 Preloaded image benchmark

| Stage / metric | Mean | Median | P95 |
| --- | --- | --- | --- |
| Preprocessing | 6.362 ms | 5.816 ms | 10.244 ms |
| Inference | 4.831 ms | 4.726 ms | 5.499 ms |
| Preloaded total (preprocess + inference) | 11.199 ms | 10.610 ms | 15.112 ms |

Script: src/benchmark_preloaded_images.py. 50 images preloaded in memory; excludes image decoding/loading. Do not compare this total directly with disk-to-prediction end-to-end timings.

# 11. Single-Image Prediction Smoke Tests

| Test | Recorded observation | Interpretation |
| --- | --- | --- |
| FP32 ONNX | Predicted Other; confidence 99.85%; preprocessing 15.190 ms; inference 5.733 ms | One single-image smoke run; timing is not statistically stable |
| FP16 ONNX | Predicted Other; confidence 99.85%; preprocessing 41.669 ms; inference 9.153 ms | One smoke run; anomalously high timings, not a fair performance comparison |

Both smoke tests indicate that the prediction script could load the corresponding model and return a class. The timings are single-run observations and should not override the repeated benchmark results. Prediction confidence is not a medical diagnosis or calibrated probability.

# 12. Hardware, Runtime and Reproducibility Context

| Component | Recorded configuration |
| --- | --- |
| Laptop CPU | AMD Ryzen 7 5800H; 8 cores / 16 logical processors |
| System memory | 15.34 GB reported |
| Laptop GPU | NVIDIA GeForce RTX 3050 Laptop GPU, 4 GB VRAM |
| Driver | 592.82; nvidia-smi reported CUDA 13.1 |
| Windows ONNX Runtime | 1.31.0; CPUExecutionProvider available |
| Windows PyTorch | 2.14.1+cpu; torch.cuda.is_available() returned False |
| Available ORT providers | AzureExecutionProvider and CPUExecutionProvider; no CUDAExecutionProvider |
| GPU inference | Not benchmarked; onnxruntime-gpu installation was not pursued after estimated large download and cancellation |
| Training GPU | Colab Tesla T4; this is separate from laptop inference hardware |

Do not report laptop GPU inference results: none were collected. CUDA reported by the NVIDIA driver does not imply the installed Python runtime has a working CUDA execution provider.

# 13. Optimization Decisions and Evidence

| Candidate / change | Evidence | Decision |
| --- | --- | --- |
| Baseline FP32 ONNX | Held-out test accuracy 92.1647%, macro-F1 91.1170%; model size about 16.05 MiB | Reference baseline |
| AMP training | Validation accuracy 93.96%, macro-F1 93.17%; duration 430.99 s, peak GPU memory 786.2 MB; no comparable baseline time | Training-time optimization completed; do not claim measured speedup |
| FP16 ONNX | Validation quality nearly unchanged; file size −49.81%; model-only CPU median −5.86%, throughput +14.14% in one run | Selected deployment candidate. Held-out test confirmed: 92.1836% accuracy, 91.1595% macro-F1, 49.81% smaller model size. See Section 18. |
| Dynamic INT8 | 72.12% validation accuracy, 65.17% macro-F1; about 4.22 MiB | Rejected due to quality degradation |
| Static INT8 | Best recorded reduced-range result 56.52% accuracy, 34.24% macro-F1; about 4.50 MiB | Rejected due to severe quality degradation |
| CPU thread tuning | Controlled median end-to-end 25.840 ms at 1 thread vs 17.423 ms at 8 threads | Use 8 threads for the measured laptop setup, subject to target-workload verification |
| Batch size 4 | 277.72 images/s vs 213.76 at batch 1 on preprocessed inputs | Potential throughput option for queued workloads; not a single-image latency improvement |
| GPU inference on laptop | No CUDA provider installed/available in ONNX Runtime | Not measured; do not claim |
| Jetson / tablet | Proposal only; no target device available for measurements | Future work only; no performance claims |

# 14. Complete Benchmark and Report Artifact Inventory

The following inventory reflects the project file names recorded during project review. JSON files remain the machine-readable source of record. A file being listed does not imply that every value inside it is reproduced in this document.

## 14.1 Benchmark JSON files

| File | Purpose / scope |
| --- | --- |
| benchmarks/batch_benchmark.json | Batch size 1/4/8/16 comparison using preprocessed images |
| benchmarks/fp32_vs_dynamic_int8_cpu.json | FP32 vs dynamic INT8 CPU comparison |
| benchmarks/fp32_vs_fp16_laptop.json | Laptop CPU FP32 vs FP16 model-only comparison |
| benchmarks/fp32_vs_static_int8_cpu.json | FP32 vs static INT8 CPU comparison |
| benchmarks/laptop_thread_benchmark.json | CPU thread-count benchmark |
| benchmarks/onnx_cpu_comparison.json | ONNX CPU benchmark results |
| benchmarks/pipeline_profile.json | Earlier pipeline profiling results |
| benchmarks/pipeline_profile_controlled.json | Older/primary controlled profile artifact |
| benchmarks/pipeline_profile_controlled_1threads.json | Controlled pipeline profile with 1 thread |
| benchmarks/pipeline_profile_controlled_4threads.json | Controlled pipeline profile with 4 threads |
| benchmarks/pipeline_profile_controlled_8threads.json | Controlled pipeline profile with 8 threads |
| benchmarks/pipeline_profile_extended_8threads.json | Extended stage profile including postprocessing and sampled RSS |
| benchmarks/repeated_prediction_8threads.json | Repeated single-image disk-to-prediction timing |

## 14.2 Report and validation files

| File | Purpose |
| --- | --- |
| reports/amp_fp32_comparison.json | Structured AMP vs baseline comparison |
| reports/amp_fp32_comparison.md | Readable AMP vs baseline comparison |
| reports/dataset_audit.json | Dataset audit |
| reports/dynamic_int8_findings.md | Dynamic INT8 findings |
| reports/fp32_vs_fp16_validation.json | FP32 vs FP16 ONNX validation comparison |
| reports/onnx_cpu_benchmark.md | Readable ONNX CPU benchmark summary |
| reports/quantization_decision.md | Rationale for rejecting quantization variants |
| reports/quantization_validation_results.json | Quantization validation results |
| reports/split_validation.json | Patient/split validation |
| reports/onnx_verification/verification_report.txt | Human-readable ONNX verification report |
| reports/onnx_verification/verification_results.json | Structured ONNX verification results |
| artifacts/baseline/test_metrics.json | Baseline held-out test metrics |
| artifacts/baseline/training_history.json | Baseline training history |
| artifacts/experiments/amp_fp16/experiment_metadata.json | AMP experiment metadata |
| artifacts/experiments/amp_fp16/training_history.json | AMP training history |
| artifacts/experiments/onnx_dynamic_int8/quantization_metadata.json | Dynamic INT8 metadata |
| artifacts/experiments/onnx_static_int8/calibration_manifest.json | Static INT8 calibration manifest |
| artifacts/experiments/onnx_static_int8/quantization_metadata.json | Static INT8 metadata |

## 14.3 Relevant scripts

| Script | Function |
| --- | --- |
| src/audit_dataset.py | Dataset audit |
| src/validate_data_splits.py | Patient-level split validation |
| src/train.py | Baseline model training |
| src/train_mixed_precision.py | AMP training experiment |
| src/benchmark_onnx.py | Model-only CPU thread benchmark |
| src/benchmark_batch_size.py | Dynamic-batch benchmark |
| src/benchmark_fp32_fp16.py | FP32 vs FP16 CPU benchmark |
| src/benchmark_preloaded_images.py | Preloaded image pipeline benchmark |
| src/profile_inference_stages.py | Earlier image-stage profile |
| src/profile_pipeline.py | Controlled loading/preprocessing/inference profile |
| src/profile_pipeline_extended.py | Extended profile with postprocessing and RSS sampling |
| src/benchmark_repeated_prediction.py | Repeated single-image timing |
| \src\evaluate_fp16_heldout_test.py | Test set evaluation of fp16 |
| src/predict_onnx.py | Single-image ONNX prediction CLI |

# 15. Source of truth and caveat

This consolidation is based on the benchmark values and artifact inventory recorded during the project work. The raw JSONs and report files in the project directory are the detailed source of record. Some INT8 benchmark JSONs and older profile JSONs may contain additional timing fields not available in the recorded notes used here; their filenames and scope are listed above rather than inventing missing values.

# 18. Final FP16 Held-Out Test Evaluation

The optimized MobileNetV3-Large FP16 ONNX model was evaluated on the isolated held-out test set containing 5,271 ultrasound images from 896 patients. The evaluation used ONNX Runtime with the CPUExecutionProvider and eight intra-op threads. The original held-out split (Train = 0) was preserved throughout; no retraining or fine-tuning was performed.

## 18.1  Overall Performance

| Metric | Value |
| --- | --- |
| Test accuracy | 92.1836 % |
| Macro-F1 score | 91.1595 % |
| Evaluation time | 16.52 minutes |
| End-to-end throughput | 5.32 images / second |

## 18.2  Per-Class Evaluation Results

All metrics are computed on the patient-isolated held-out test split. Support counts reflect the class distribution in that split.

| Class | Precision | Recall | F1-score | Support |
| --- | --- | --- | --- | --- |
| Other | 0.91 | 0.85 | 0.88 | 1,612 |
| Fetal brain | 0.99 | 0.98 | 0.99 | 1,472 |
| Fetal thorax | 0.85 | 0.96 | 0.90 | 660 |
| Maternal cervix | 0.99 | 0.99 | 0.99 | 645 |
| Fetal femur | 0.84 | 0.88 | 0.86 | 524 |
| Fetal abdomen | 0.84 | 0.87 | 0.85 | 358 |
| Macro average | 0.90 | 0.92 | 0.91 | 5,271 |
| Weighted average | 0.92 | 0.92 | 0.92 | 5,271 |

## 18.3  Comparison with the FP32 Baseline

| Metric | FP32 Baseline | FP16 Optimized |
| --- | --- | --- |
| Held-out test accuracy | 92.1647 % | 92.1836 % |
| Held-out test macro-F1 | 91.1170 % | 91.1595 % |
| ONNX model size | 16.05 MiB | 8.05 MiB |
| Model size reduction | — | 49.81 % |

The FP16 ONNX model reduced model size by approximately 49.81 % while maintaining comparable held-out test accuracy and macro-F1. The small metric differences (< 0.02 pp accuracy, < 0.05 pp F1) should not be interpreted as a statistically significant accuracy improvement.

Note: The reported throughput figure (5.32 images/s) includes disk I/O, preprocessing, and model inference and should not be directly compared with the model-only inference benchmark results in Section 6.

## 18.4  Section Conclusion

Accuracy: 92.18 % on 5,271 held-out images (patient-isolated split).

Macro-F1: 91.16 %, meeting the deployment quality threshold.

Model size: 8.05 MiB -- 49.81 % smaller than the FP32 ONNX baseline (16.05 MiB).

Generalization: Accuracy and F1 are consistent with the FP32 baseline; no degradation from FP16 conversion.

Deployment: FP16 ONNX is the selected post-training optimization for this CPU-based deployment experiment.

Future work: Performance on Jetson Orin Nano and tablet hardware has not been measured and is documented as a future deployment extension.